Confocal Tutorial

- Follow instructions for turning on components

- Sign on as “nelson” password, “funfun” – enter into Fluoview program

- Put on 60X oil objective, add drop of oil directly to objective, match the objective to fluoview setting on the screen (60X)

- Flip down the light source

- Turn dial (on right of scope) to eye symbol

- Put the slide on the stage

- Bring objective up to the slide – start over if introduction of bubble

- Light on by pushing button on front left of scope and by pulling out shutter on top of scope – lamp housing

- Focus on cells and then clip into place

- Turn fluorescent light on – turn off brightfield = light button off, close shutter at top of scope, turn cube on right of scope until blue light emission from the scope (MF filter)

- Focus until on cell for picture capturing

- Move laser intensity if needed (20 for weak signal, 6 for strong signal)

- To take a picture – Move filter cube back to brightfield setting, move eye symbol on dial to SP, close shutter at top of scope

- On Fluoview program – hit “focus” at top of screen

- On FITC channel - Turn up the gain all the way, turn down the PMT most of the way

- On RGB bar at bottom of screen, click HighLow button and then increase the PMT button until the screen shows gray with red dots bleeding through

- Change the RGB image back to Green

- Focus on the object – might need to zoom in/out or move around

- Move PMT until looks good

- Scan to take the picture

- Save by right clicking on the image – save to “other” – “nelson” – make new folder with “N.Robinson” and save to that

- To take slice picture (for 3-D recreation and movies) – click on the sidebar “Zstage” (previously on “Scan”),  hit XYZ (previously on XY)

- Set the limits by increasing with arrows, hit “set” and then decreasing, hit “set”  to the upper and lower limits of the cell

- Set the step size (use b/n 0.5 – 0.8um steps)

- Hit “XYZ” button

- When finished, hit “Series done”

- Save slices picture

- Hit “Locked” checkmark twice to unlock the fine focus on the microscope

- Play when finished or recreate 3-D by clicking at the top folded bar on “Tile” to advance to the right – click until “Visualize” – use Live vs. 3-D

- To save to network – on desktop – go to Entire Network – bsd-dean – files – type “bsd-dean\ncampbel” password “*******” – dnelsonlab – n.robinson