## Slide 1: Flow Basics
Cytometry and Antibody Technology Service Facility
University of Chicago, 2021
A Self-Defense Guide to Flow Cytometry

## Slide 2: A flow cytometer is a bad microscope
Intrinsic information
Scattered light
Forward Scatter (FSC)
  Relative size
Side Scatter (SSC)
  Granularity
Extrinsic information
Emitted light from fluorophores/dyes
Antibody against proteins of interest
Dyes binding to a specific part of the cells

## Slide 3: The challenge of extrinsic parameters: Resolving signals from cell autofluorescence
Autofluorescence
(no fluorophore)
Fluorophore
(antibody or dye)
Maximizing resolution
  Marker expression
  Fluorophore brightness
  Instrument settings
Signal Specificity

## Slide 4: Fluorpophores and Dyes: Excitation, Emission and Brightness
Excitation
Emission
Learn more about Fluorophores in this episode of the ChUG Podcast

## Slide 5: A look inside the box
Fluidic system
Optical layout
Electronics
Sheath Tank
Flow Cell
Laser
Detection System
Waste Tank
Analysis

## Slide 6: Fluidic system
Sheath
Tank
Waste
Tank
Line Pressure
Vacuum
Sample
Pressure
(Variable)
Sheath Pressure
(Constant)
Sample
Tube

## Slide 7: The Flow Cell
In order to get information about each cell individually, we need cells to pass the interrogation point one at a time
The hydrodynamic focusing within the flow cell aligns the cells

## Slide 8: Sample Differential
Increasing the speed widens the sample core
Results in unequal illumination of cells and increase in coincidence

## Slide 9: Optical bench layout
Intrinsic parameters
  Light Scatter
Lasers also excite fluorophores and dyes
  collected by a detector array

## Slide 10: Light Scatter
Lymphocytes
RBC, debris, dead cells
Monocytes
Granulocytes
FSC
SSC

## Slide 11: Traditional Flow Cytometers
505LP
555LP
640SP
488/10
530/30
585/42
695/40
SSC
FITC
PerCP Cy5.5
PE
FSC Detector

## Slide 12: Spectral Flow Cytometry
Never have to worry about filter sets again
Separate fluorophores with similar emission
Manage larger number of markers more efficiently
Autofluorescence extraction
Trad flow
spectral flow

## Slide 13: Multi-laser instruments and pinholes
Detector Path laser 1
Detector Path laser 2
Increase the number of usable markers
Reduce the amount of spillover between some combination of fluorophores

## Slide 14: Various types of detectors used in different instruments
https://hamamatsu.magnet.fsu.edu/articles/photomultipliers.html
https://commons.wikimedia.org/wiki/File:APD3_German-ru.svg
PMT
APD
siPM
Here is a review of the main detectors used in flow

## Slide 15: Threshold
Time
Signal intensity

## Slide 16: The Voltage Pulse
Time
Signal intensity

## Slide 17: Measurement of the pulse
Signal Intensity
Time
Signal Intensity
Width
Area
Height

## Slide 18: FCS Data File
Flow Cytometry Standard (FCS)
  Accepted Standard in flow cytometry world
  Usable on third party analysis software packages
Listmode File containing
  All measurements
  Instrument settings acquisition information
  Keywords

## Slide 19: Flexibility in data presentation
Graphic analysis (typical)
Unsupervised clustering (high parameters)
Modelization

## Slide 20: Panel Design
Choosing Fluorophores
Choosing Instrument

## Slide 21: Good panel design
Great panel design aims at improving the resolution of your different markers
Variables include
  Biology of the markers
  Brightness of Fluorophore
  Spectral Overlap
Check the Sample preparation guide for more information

## Slide 22: Overlapping emission generates artefact that needs to be corrected
APC
AF700
APC stained cells

## Slide 23: Spillover increases error spread

## Slide 24: Tandem Fluorophores are great but noisy
Energy transfer by FRET but:
  Donor molecule also emits photons
  Acceptor molecule is excited by lasers as well
Tandem dyes go bad over time

## Slide 25: Easy start for a simple panel
Fluorophore | Laser
BV421 | 405nm
FITC | 488nm
PE | 561nm
APC | 640nm
Excited by different lasers
Non-overlapping emission
Single molecule (non-tandem)

## Slide 26: Benchtop instruments variations
Usable fluorophores
  Number of lasers
  Number of measurable parameters
Features available
  Spectral unmixing
  Acoustic Focusing
  Sample Loader
  Imaging

## Slide 27: Traditional Flow Cytometers
LSRII/Fortessa
  Plate loader (1 - screening)
Attune NxT
  Acoustic Focusing
  True count
  Plate Loader
Quanteon/Penteon
  Plate Loader
  True count
  High Dynamic Range

## Slide 28: Spectral Flow Cytometers Cytek Aurora
Measures the entire emission of each fluorophore and unmix the signal in discrete parameter measurement
Can be used as a traditional flow cytometer
Not limited by filter selection

## Slide 29: Cell sorters
Flow cytometers that allows for the collection of one or multiple sub-population of cells
Droplet-based or Mechanical sorting
Staff operated in general, but training available

## Slide 30: Noteworthy platforms
Image Stream MarkII
  Takes pictures of each event in up to 12 channels
Multiplexing
https://www.thermofisher.com/blog/behindthebench/luminex-bead-based-immunoassays-drive-immunoassays-towards-higher-content-biomarker-discovery/
Luminex 200
Bead-based assay
30+ analytes
MSD
ECL
Up to 10 analytes

## Slide 31: Getting started
Create plots that will give you meaningful information
  FSC x SSC
  Fluorescence
Histograms are not as informative as one would like them to be, go with dot plots whenever possible

## Slide 32: Lin vs. Log
FSC is usually displayed on Lin scale
SSC can be displayed on both linear or logarithmic scale
Fluorescent parameters are displayed on log scale except for specific applications (DNA content analysis, …)
SSC on Log
SSC on Lin
Flow cytometry data will be displayed on linear or logarithmically scaled axes
  Log scale allows higher dynamic range

## Slide 33: Setting up voltages on FSC vs. SSC
FSC
SSC
Voltage
Bring the population of interest in the center of the graph
  Provides optimal resolution of population
  Avoid giving too much importance to debris/losing cells off-scale

## Slide 34: Setting up threshold
FSC
SSC
Threshold
Avoid including too much debris
Make sure you are not loosing information

## Slide 35: Fluorophore detectors voltage
PE
SSC
Voltage
Voltage needs to be high enough to allow for resolution of the positive and negative fractions
Voltage set too high does not provide better resolution but increase noise in negative fraction

## Slide 36: Default instrument values are optimized
The instruments have been characterized and the optimal voltage/gain values are set in the default experiment settings
You may need to lower voltage / gain to avoid strutting the detector
You don’t win anything by increasing the values

## Slide 37: Exercise: How do these fluorophore voltages look?
1A
1B
2A
2B
3A
3B
4A
4B

## Slide 38: Dealing with spillover
Alexa 488
PE
Uncorrected
Corrected
Signal correction in multi-parameter experiment made necessary by fluorophore spectral overlap
Compensation
  Substraction of a percentage of the signal of a fluorophore (that happens to be bleeding in another detector) from the measured light in that detector
Unmixing
  Calculate the contribution of each known fluorophore’s spectra to the total collected emission signal

## Slide 39: Do I need compensation controls? What strategy did you use?
Avoidance
Pick fluorophores which will not overlap with one another
No need for controls
Overlap included
If spillover can’t be avoided, correction is needed
Prepare single-stained controls
Notes: Something about how the data should look like in the end

## Slide 40: Spillover controls
Compensation controlrules
  Must have a positive / negative fraction
  The positive and negative fraction must have the same baseline autofluorescence
  Positive fraction’s brightness similar or higher than sample
  Use the same fluorophores / dyes for sample and controls
  Save a good number of events for both positive and negative fraction
Use of software recommended (necessary in the case of unmixing)
You will need to run your set of controls and compensate on each new experiment
More information on compensation on our website

## Slide 41: Data Analysis
Cleaning Data
Gates justification

## Slide 42: Cleaning Up Data I
Removing Doublets
Removing Acquisition Mishap

## Slide 43: Cleaning Up Data II
Removing Dead Cells

## Slide 44: Deciding on gates: FMO Controls
FMO PE
Sample

## Slide 45: Deciding on gates: Backgating

## Slide 46: Resources
CAT Facility website
ChUG Website

## Slide 47: Acknowledgment
Purdue University Cytometry Laboratories website: http://www.cyto.purdue.edu/
  Dr. Robert Murphy, Carnegie Mellon University- Basic Theory 1 and 2 powerpoint slides
Sydney Cytometry Facility tutorial: The impact of adjusting PMT voltages on spillover and compensation
BioLegend spectraviewer
ThermoFisher spectraviewer
ThermoFisher website training material